## Slide 1: OscilLometric Blood Pressure
Miodrag Bolic

## Slide 2
Blood pressure definitions

## Slide 3
Definitions
Auscultatory method is an indirect method to measure arterial blood pressure that relies on the detection of Korotkoff sounds under an occluding cuff.
Electronic sphygmomanometer is an instrument used for the indirect (noninvasive) measurement of arterial blood pressure that automatically determines blood pressure from signals obtained from the cuff and/or a separate transducer.
Oscillometry is an indirect method to measure arterial blood pressure that relies on the measurement of oscillations in an occluding cuff.
Non-invasive blood pressure measurements (NIBP) is normally any cuff-based method for measuring BP
Ambulatory BP monitoring (ABPM) is a noninvasive, fully automated technique in which BP is recorded over an extended period of time, typically 24 hours.
Home blood pressure monitoring (HBPM) that typically involves measurements over a longer periods of time, in stationary position while being awake.

## Slide 4
Measuring blood pressure
  S. Tavoularis, Measurement in Fluid Mechanics, New York : Cambridge University Press, 2005
    One can measure pressure of a fluid inside of a pipe only if a sensor is inserted in it!
  Direct measurement of blood pressure is invasive, and, as such, it has a very limited clinical value.
  Non-invasive blood pressure measurement is the only application where direct measurement methods from other fields could not be applied for practical and ethical reasons.
  BP cannot be measured non-invasively but only estimated from indirect measurements (Korotkov sounds, cuff pressure oscillations, tonometry, etc)
  It is the result of internal REGULATION

## Slide 5
Variability of blood pressure
157/71
SBP/DBP
169/63
SBP/DBP
Up to 20 mmHg
Cardiovascular
Mayer Waves
Respiration

## Slide 6
Oscillometric method
A cuff is connected to a subject
A cuff is inflated to a point that the blood does not flow through the artery any more
The pressure is slowly released
As the blood start flowing, the cardiac pulses are superimposed to the cuff pressure
The pressure is measured and oscillometric algorithm is applied to estimate systolic and diastolic pressure

## Slide 7: Age & Cuff Deflation Waveform
Pulse waveform at different pressure
Pulse waveforms characteristics change at different pressure points in oscillometric cuff deflation pressure waveform

## Slide 8
How can we process the pulse
Extracting features from a single pulse
Looking at the function: the feature vs. time or vs. pressure
Estimating the values of Systolic and Diastolic from the function.
Feature
Feature
Time or
Cuff Pressure
Preprocessing
and
processing
SYS,
DIA,
MAP
Notes: Pulse waveform is a result of complex interaction between the left ventricle and blood circulation system. Pressure signal usually contains more information than is captured BP values and most of this information can be extracted from the pulse morphology,
Pulse represents 1 cycle of arterial blood pressure
Pulse is a graph of pressure over time (figure)
Pulse has first(systolic) peak, and sometime the second peak

## Slide 9
How can we process the pulse – Innovation?
What features to extract that would lead to determination of SYS, DIA and MAP
  that have physiological sense.
  that have correlation with age, some conditions,…
Feature
Feature
Time or
Cuff Pressure
Preprocessing
and
processing
SYS,
DIA,
MAP
Notes: Pulse waveform is a result of complex interaction between the left ventricle and blood circulation system. Pressure signal usually contains more information than is captured BP values and most of this information can be extracted from the pulse morphology,
Pulse represents 1 cycle of arterial blood pressure
Pulse is a graph of pressure over time (figure)
Pulse has first(systolic) peak, and sometime the second peak

## Slide 10
Pulse processing – Amplitude based
Maximum amplitude algorithm
Feature of interest: Maximum, area or amplitude of pulses
Physiology: maximum of the function is found to correlate with MAP
Preprocessing
  Find maximum of pulses
  Generate the function -envelope
  Clean the envelope
Processing
  MAP - pressure that corresponds to the maximum
  SYS = MAP * K_SBP
  DIA = MAP * K_DBP

## Slide 11
Oscillometric algorithm

## Slide 12
Pulse processing – Amplitude based
Disadvantages
Disadvantages:
Does not use the wealth of information from the pulse
It is based on empirically derived coefficients for computing systolic and diastolic BP
Solution:
Estimate coefficients
Coefficient-free processing
  Neural networks
  Analysis of the pulse

## Slide 13
Envelope processing – Amplitude based
Estimating coefficients
Goal: Adjust coefficient of the oscillometric algorithm
Method: Include personal information in the model
M. James, Simpliﬁed Model for the Design of an Oscillometric Blood Pressure Measuring System, PhD Thesis, University of Guelph, 2012.
Notes: paper from guelph

## Slide 14
Envelope processing – Amplitude based
Neural Networks
Neural networks (NNs) can approximate almost any nonlinear relationship that exist between inputs and outputs
Existing work:
  The raw oscillometric waveform envelope (OMWE) is evenly sampled at specific increments of CP
  The resultant samples are fed to the NN as input.
Our approach:
  Modeling of the oscillometric waveform envelope as sum of two Gaussian functions as suggested in the literature
  Parameters from Gaussian functions are extracted
  These parameters are input to the network

## Slide 15
Envelope processing
Neural Networks
Preprocessing
Processing

## Slide 16
Features that can be extracted from a single pulse and tracked in time
Systolic & diastolic amplitude
Systolic, diastolic slope
ΔT/T ratio
Area under systolic curve
Pulse duration
Existence of dicrotic notch
Stiffness Index (SI)
Reflection Index (RI)
Feature
Feature
Time or
Cuff Pressure
Preprocessing
and
processing
SYS,
DIA,
MAP
Notes: Pulse waveform is a result of complex interaction between the left ventricle and blood circulation system. Pressure signal usually contains more information than is captured BP values and most of this information can be extracted from the pulse morphology,
Pulse represents 1 cycle of arterial blood pressure
Pulse is a graph of pressure over time (figure)
Pulse has first(systolic) peak, and sometime the second peak

## Slide 17
Analysis of the pulse
Other features
diastolic
MAP
Reflection Index
)
Feature
Processing:
MAP – maximum of the function

## Slide 18
Modeling
Accuracy needs to be improved
  Rely on scientific approach - modeling
Applications
  Development of new algorithms
  Understanding the mechanisms
  Setting parameters of existing algorithms
  Calibration
Notes: File: model and measurement

## Slide 19
Innovation - modeling
Sum of harmonically related sinusoids modulated by respiration
Arterial pressure-area relationship?
Oscillometric pulses are proportional to the oscillations of the arterial lumen area
Notes: File: model and measurement

## Slide 20
Innovation - modeling
Notes: File: model and measurement

## Slide 21
Oscillometric device Hardware
Notes: File: model and measurement

## Slide 22
Oscillometric device - Control
Notes: File: model and measurement

## Slide 23
Data acquisition

## Slide 24
Subject and Measurements
Adherence to measurement recommendations
Recommendations
  Patients must remain silent during measurements,
  be seated correctly with back support and legs uncrossed,
  must have rested at least 5 minutes prior to taking the measurement.
  should reside in a quiet environment.
Problem
  current state-of-the­ art BP devices are not capable of sensing incorrect usage
  Since only measurements following the recommendations are considered reliable
  data from the reported studies could be indeterminate

## Slide 25
Factors that Can Elevate Blood
Pressure Readings
http://www.suntechmed.com/

## Slide 26
Subject and Measurements
Adherence to measurement recommendations
A number of sensors is used
A number of research papers have appeared recently

## Slide 27
Subject and Measurements
Adherence to measurement recommendations
The system monitors activity and posture
The system records audio data and classifies as speech or silence.
S. Wagner et. al. , Context Classification during Blood Pressure Self-Measurement using the Sensor Seat and the
Audio Classification Device, PervasiveHealth Workshop, 2012.

## Slide 28
Subject and Measurements
Adherence to measurement recommendations